510(k) Summary

BreatheStim, 5270 California Avenue, Irvine, California 92617

BreatheStim Appliance

Intraoral Device for Snoring and Mild Sleep Apnea Treatment

Device, anti-snoring, mild sleep apnea 
(Class II, 21 CFR 872.5570, Product Code LRK)

The BreatheStim Anti-Snoring and Mild Sleep Apnea Appliance is an integrated system for intraoral use. The  device is a combination bioelectric stimulation microcurrent and elastic mouthpiece. It features elastic elements that span the device that are designed to keep the tongue in place during sleep, instead of blocking the airway and contributing to snoring and to reduce inflammation and tighten tissues to provide long lasting breathing relief including elimination or reduction of snoring. The product is non-sterile and provided in a sealed box with directions for use.

Indications for Use

BreatheStim is intended for use by adult patients (18 years or older) as an aid to reduce snoring and to provide relief of mild sleep apnea.

10. Technological Characteristics

The appliance is equivalent in design and functionality to the predicate, prescription Dr. Greenbergs Hybrid device (K111680). Both devices work by depressing the tongue during sleep to open the airway and reduce the incidence of snoring, in addition to advancing the mandible.  Both mouthpieces use the similar materials and manufacturing processes. Both devices are provided non-sterile. The primary difference between the subject and predicate device is the addition of faint microcurrent signals to aide in reducing inflammation and to tighten tissues.

The appliance has similar labeling as the OTC reference predicate devices, SnoreRx and ZQuiet. It utilizes a validated questionnaire to assess the risk of sleep apnea, and contains appropriate warnings and labeling for the device to be used without a prescription.

|  | Subject | Prescription Predicate | OTC Reference Predicate | OTC Reference Predicate |
| --- | --- | --- | --- | --- |
| Name | BreatheStim Anti- Snoring and Mild Sleep Apnea Appliance | Dr. Greenburg’s Hybrid | SnoreRx | ZQuiet |
| Classification Name | Device, anti- snoring. | Device, anti- snoring | Device, anti- snoring | Device, anti- snoring |
| 510(k) Number | Pending | K111680 | K170825 | K180124 |
| Class | II | II | II | II |
| Product Code | LRK | LRK | LRK | LRK |
| 21 CFR | 872.5570 | 872.5570 | 872.5570 | 872.5570 |
| Technology & Mechanism of Action | Repositioning the tongue, mandibular advancement | Repositioning the tongue, mandibular advancement | Mandibular advancement | Mandibular advancement |
| Indications for use | BreatheStim is intended for use by adult patients (18 years or older) as an aid to reduce snoring. | This appliance is indicated for persons 18 years or older, who wish to reduce the incidence of snoring and/or mild to moderate sleep apnea. | “SnoreRx is intended for use on adult patients 18 years of age or older as an aid for the reduction of snoring” | “ZQuiet is intended as an aid in the reduction of snoring for adults at least 18 years old.” |
| Prescription Status | OTC | Prescription | OTC | OTC |

5-1

| Provided Sterile | No | No | No | No |
| --- | --- | --- | --- | --- |
| Materials | Hard outer shell with thermoplastic filling material and elastic band | Hard outer shell with thermoplastic filling material and elastic band |  | Thermoplastic Elastomer with Colorant |

10.1 Technology – Mechanism of Action

The subject device and prescription predicate device (K111680) are identical in mechanism of action with the exception of mild microcurrent signaling added into the BreatheStim device. They both work by utilizing an intraoral mouthpiece to reposition the tongue in conjunction with a slight advancement of the mandible.  The microcurrent signaling is designed to primarily reduce risk of relapse after the mechanical device is removed.

The mechanism of action for the other two OTC predicate devices, the SnoreRx and ZQuiet, is solely advancing the mandible. With any intraoral device that alters the mandible position, mandibular changes can lead to pain or discomfort in the temporomandibular joint (TMJ). The subject device has no new risks in regards to mandibular advancement.

Utilizing the same mechanism of action as the prescription predicate the subject device raises no new concerns regarding effectiveness or safety as compared to predicate devices.

10.2 Technology – Indication for Use

The subject device uses language for the indication for use that is identical in substance to previously- approved OTC devices for the reduction of snoring. The subject device is “intended for use by adult patients (18 years or older) as an aid to reduce snoring.”

This is almost identical to the indications for use for the two OTC predicate devices, the SnoreRx and ZQuiet, whose indications for use are “SnoreRx is intended for use on adult patients 18 years of age or older as an aid for the reduction of snoring” and “ZQuiet is intended as an aid in the reduction of snoring for adults at least 18 years old,” respectively.

As the indications for use are identical, there is no new concerns regarding effectiveness or safety as compared to predicate devices.

10.4 Technology - Prescription Status

The subject device is intended to be sold over-the-counter. Product labeling has therefore eliminated prescribing information. Two of the predicate devices, SnoreRx and ZQuiet, have clearance for over-the- counter use.

Warnings and contraindications are substantially equivalent to those of over-the-counter predicate devices. The labeling specifically states that the device does not treat obstructive sleep apnea, and includes a   validated questionnaire is included in the labeling to help mitigate the risk of undiagnosed OSA among potential BreatheStim patients.

The similar predicate prescription device (Dr. Greenburg’s Hybrid, K111680) has been on the market for more than ten years. Due to its postmarket record of substantially equivalent performance to other devices of this type and substantially equivalent labeling and instructions for use to that of the over-the-counter predicate devices, there are no new concerns regarding effectiveness or safety as compared to predicate devices.

10.5 Technology - Materials

To support this application, the sponsor has included Biocompatibility testing results demonstrating the subject device is biocompatible under ISO 10993, including cytotoxicity, sensitization, and irritation testing. The device was shown to meet the requirements of the ISO 10993 guidelines for each of these tests. Other biocompatibility tests recommended for permanent mucosal contact were not conducted because all materials used in the device are certified USP Class VI and used in many other currently marketed medical devices.

Because the same materials and manufacturing processes are used as the currently marketed predicate device, there are no new concerns regarding effectiveness or safety as compared to the predicate device.

11. Non-Clinical Testing

The proposed BreatheStim OTC device was compared to across all design and performance charactersics of the cleared predicate in bench testing. It was found in this testing to be substantially equivalent predicate device.

The sponsor conducted a risk analysis on the device in accordance with ISO 14971:2007, taking into account the issues raised in the FDA Guidance Document “Class II Special Controls Guidance Document: Intraoral Devices for Snoring and/or Obstructive Sleep Apnea – Guidance for Industry and FDA.” All identified risks have been addressed through device design or with communication with the user through the instructions for use.

See Addendum A for a summary of non-clinical testing and our full risk analysis table.

12. Biocompatibility

The materials and manufacturing process are similar to the currently marketed materials in the prescription predicate device.

Testing results that have been included include cytotoxicity (ISO 10993-5), sensitization (ISO 10993-10) and irritation (ISO 10993-10). As all materials used in the device are certified USP Class VI and used in many other currently marketed medical devices, no additional biocompatibility testing was conducted.

Class VI certification is less stringent than 10993

13.  Stimulator Testing

The Mettler Electronics of Anaheim, California stimulator used by BreatheStim has FDA 510 K11017 June 26th, 2012 market clearance for improving circulation, muscle re-education, reduction of muscle atrophy and mild pain relief all applicable to this application of use.

Full Spec Sheet > https://www.mettlerelectronics.com/images/products/neuromusc_stim/240_specifications.pdf

Mettler 240 Instruction Manual > https://mettlerelectronics.com/wp-content/uploads/240man.pdf

Performance Testing: The Sys*Stim 240 has been certified for:

" ITS certification for ETL and C-ETL Listing

* Compliance with EMC, EN 60601-1 -2: 2001 and FCC's 47 CFR,

FCC Sub Part 18

* Compliance to IEC 60601-2-22: 2007

* Compliance to IEC 60825-1: 2007

* Compliance to IEC 62471: 2006

- Compliance to EN60601-2-10, (IEC 601-2-1 0) 2001

The Sys*Stim 240 has an optional battery pack so that you can take either electrical stimulation or light therapy to patient. A carrying case is also available which holds the units and all the accessories necessary for therapy on the road.

The Sys*Stim 240 has been certified by Intertek Testing Services to meet the requirements for ETL Listing per the following standards:

UL 60601-1 Standard for Safety Medical Electrical Equipment, Part 1: General Requirements for Safety.

CSA C22.2.601.1 M90 Medical Electrical Equipment – Part 1 General Requirements for Safety

IEC / EN 60601-1 Standard for Safety Medical Electrical Equipment, Part 1: General Requirements for Safety

IEC / EN 60601-2-10 Medical Electrical Equipment Part 2: Particular Requirements for the Safety of Nerve and Muscle Stimulators

5

Mettler Electronics Corp. — Rev.I_09/23/19

IEC / EN 60601-2-22 Medical Electrical Equipment – Part 2-22: Particular Requirements for Safety and Essential Performance of Surgical, Cosmetic, Therapeutic and Diagnostic Laser Equipment

IEC / EN 60825-1 Safety of Laser Products Part 1: Equipment Classification and Requirement

IEC / EN 62471 Photobiological Safety of Lamps and Lamp Systems

In addition, the Sys*Stim 240 also meets the following standards for radio frequency emissions and immunity:

IEC / EN 60601-1-2 Medical Electrical Equipment—Part 1-2: General Requirements for safety—Collateral Standard Electromagnetic Compatibility Requirements and Test

FCC 47 CFR, FCC Sub Part 18 Industrial, Scientific and Medical (ISM) Equipment

Mettler Electronics Corp. has been certified by VTT Expert Services LTD to be compliant with EN ISO 13485:2003 and MDD 93/42/EEC Annex II requirements. In addition, Mettler is certified by DQS Medizinprodukte GMBH to be compliant with ISO 13485:2003 (CMDCAS) Canadian Medical Device requirements.

The new Mettler 240L developed in collaboration with Leonhardt Ventures LLC and proprietary to Leonhardt Ventures LLC is currently under FDA review which is the same circuity of the original 240 but miniaturized to hand held palm size share and with firmware added to safety protections for at home use.  Leonhardt Ventures LLC and Mettler Electronics have granted an exclusive license to BreatheStim for the use of the 240L portable stimulator for the indication of use of anti-snoring and sleep apnea treatment.

14. Clinical Testing

Clinical testing was performed in association with this submission in 10 patients at a breathing disorder clinic in Irvine, California working with Dr. Blair Bittner.  We have included in the attached addendum B photos and data from all 10 of these cases.  No safety issues arose in these studies. All 10 patients had relief of snoring and had improving night time and day time breathing after treatment many of them exhibiting substantial improvement.

15. Conclusions

The results of the comparison of design, materials, intended use, labeling, and technological characteristics demonstrate the subject device is substantially equivalent in safety and efficacy the legally marketed predicate devices. Therefore, the sponsor concludes the proposed BreatheStim device is substantially equivalent to the identified predicate devices.